Type 2 F and B checklist (2BS, 2BL1, 2BL2)

DTMROC Inputs:

Starting from Pin # 1, clockwise:

TR14 – 15 P,N: 			threshold scan

tp_bias				think it’s fine

tp_even, TP_odd			test pulse scans

th_TR0, th_TR1			high threshold scan….have verified for a few chips

th_D0, th_D1				threshold scans

Config_select				goes to PADJ1, PADJ2 – Mitch checked

Shaper_select				goes to Xel – Mitch checked

ASDBLRpwrsense			tested

ASDBLRpwrjmp			tested

SpareInpsense			tested

spareInpjmp				tested

enable_decoup2			Mitch is not sure of the status; not important

cmd_in_neg, cmd_in_pos		threshold scan

bc_neg, bc_pos			threshold scan

hard_reset_B_neg,….pos		have checked on scope, polarity has to be correct to work

cmd_out_pos, com_out_neg		register readout

data_out_net, data_out_pos		threshold scan

enable_decoup1			Mitch is not sure of  the status; not important

TRST,TCK,TDI,TMS,TD0		NC (JTAG) (needs reviewing for Bjorn’s boards)

Address 0-5	Known to work on AR1F & AR1B – connected & way we want them.

TR00 – 01 P,N			threshold scan

TR02 – 13 P,N			threshold scan

NAIS Connector to Paddle Card

Do above tests test all signals?

ASDBLR

Power: VCP,VCS,VES,VCD,	checked by basic operation

VED,VEDR

Ternary Inputs: 			ohm-meter test or noise tests on the module

Ternary Outputs:			threshold scan

PVCDS				stuffed for what current?

PTH_D				threshold scan

PTH_TR				test-pulse scan at high threshold

XEL, PADJS1,PADJS2		Mitch looked at

BLBIAS				NC

TST_O				test pulse

TST_E				test pulse

PEN_SH, PEN_BL			Mitch

MON_A1,MON_B1

MON_A8, MON_B8			OK – not normally used

Board Level:

Noise Tests: performed for four AR1BS & AR1BL boards, and several AR1FL and AR1FS boards on the module. Verified that which Vdd caps should be removed on AR1FL boards and that copper tape (or equivalent) should be added on AR1FS. Whether or not to remove caps on AR1B boards remains to be checked…Do we wait for this for AR1B assembly?

Is the termination correct: have looked at the receiving termination , about 100 ohms. We’ve never investigated the back termination resistors. Currently it is 120 ohms. Proabably OK. But Rick raises issue of size of signal at the PP. Mitch says we see 100+ mV but Rick raises the question of what length of cable. Mitch – we should test “as stuffed”. Rick would like to review..what is the longest cable we have ever used.

Value of resistors between analog and digital ground isn’t critical. Ben verified that the resistors are all 10 ohms.

Test Pulse – Mitch is happy with the filter components on barrel boards, but would really like to see the response when we try to try high threshold test.

Report from Ben: only two problems on all 8 ARB boards. One shorted input, couldn’t locate exactly where it was; one channel which gets much noisier during test pulse.

Would see a 300kHz threshold of 250 or so where most all channels are around 180 (this for an injected charge of 10 DAQ cnts). NOTE: they looked at three other boards and didn’t see anything.

So 2 channels out of 1300.

Found a few more channels with large offsets for Gabe to check against his measurements.

Questions:

Ben: do we have a final plan for labeling these boards?

Cooling Plate Mounting – DONE in terms of mates OK and noise measurements. Temperature measurements haven’t yet been repeated.